## Slide 1
Uncertainties in measurement
Mass fraction calculation options
Recalculations using new wt% and Ice cp measurements
Comparisons with SeaFreeze EOS and Tillner-Roth & Friend’s EOS
NH3 EOS errors of rho and cp w.r.t T and m
Periodicity in melting phase
ADH double peaks
Summary

## Slide 2
Ice Cp: are we concerned with the difference between
  0.25 K/min and 0.1 K/min
  Our experiment and Feistel and Wagner
  0.1 K/min variability at the start and in the middle
Mass Fraction Determination: if based on measurements
  Density is inconsistent:
Questions

## Slide 3
From laboratory measurements
Based on density, volume and concentration of NH4OH, with mass of sample
Mass fraction calculation option 1
No. | (g/ml) | (ml) | (wt%) | (g) | (wt%)
1 | 0.789 ± 0.005 | 0.44 ± 0.005 | 20 ± ? | 4.8873 ± 0.001 | 1.42 ± 0.03
2 | 0.789 ± 0.005 | 0.87 ± 0.006 | 20 ± ? | 4.8402 ± 0.001 | 2.84 ± 0.04
3 | 0.789 ± 0.005 | 1.50 ± 0.01 | 20 ± ? | 4.6293 ± 0.001 | 5.11 ± 0.07
4 | 0.789 ± 0.005 | 2.00 ± 0.01 | 20 ± ? | 4.5853 ± 0.001 | 6.88 ± 0.09
5 | 0.789 ± 0.005 | 2.50 ± 0.02 | 20 ± ? | 4.5202 ± 0.001 | 8.7 ± 0.1
No. |  |  |  |  | 
1 | 0.789 ± 0.005 | 0.44 ± 0.005 | 20 ± ? | 4.8873 ± 0.001 | 1.42 ± 0.03
2 | 0.789 ± 0.005 | 0.87 ± 0.006 | 20 ± ? | 4.8402 ± 0.001 | 2.84 ± 0.04
3 | 0.789 ± 0.005 | 1.50 ± 0.01 | 20 ± ? | 4.6293 ± 0.001 | 5.11 ± 0.07
4 | 0.789 ± 0.005 | 2.00 ± 0.01 | 20 ± ? | 4.5853 ± 0.001 | 6.88 ± 0.09
5 | 0.789 ± 0.005 | 2.50 ± 0.02 | 20 ± ? | 4.5202 ± 0.001 | 8.7 ± 0.1

## Slide 4
From using the temperature of the measured liquidus peak to determine theoretical mass fraction via the liquidus curve
Mass fraction calculation option 2
Liquidus Peak T (K) | (wt%) (option 2) | (wt%) (option 1)
274.5 ± 0.5 | ? | 1.42 ± 0.03
271.0 ± 0.5 | 2.5 ± 0.5 | 2.84 ± 0.04
268.2 ± 0.5 | 4.6 ± 0.7 | 5.11 ± 0.07
262.9 ± 0.5 | 7.9 ± 0.6 | 6.88 ± 0.09
259.6 ± 0.5 | 9.7 ± 0.5 | 8.7 ± 0.1
207.2 ± 0.5 | 27.0 ± 0.2 | 28 – 30
Liquidus Peak T (K) |  | 
274.5 ± 0.5 | ? | 1.42 ± 0.03
271.0 ± 0.5 | 2.5 ± 0.5 | 2.84 ± 0.04
268.2 ± 0.5 | 4.6 ± 0.7 | 5.11 ± 0.07
262.9 ± 0.5 | 7.9 ± 0.6 | 6.88 ± 0.09
259.6 ± 0.5 | 9.7 ± 0.5 | 8.7 ± 0.1
207.2 ± 0.5 | 27.0 ± 0.2 | 28 – 30

## Slide 5
From NH3 EOS, extract mass fraction based on density of sample
Assumed measurements done at 25°C
Mass fraction calculation option 3
No. | (g/ml) | Molality (mol/kg) | (wt%) (option 3) | (wt%) (option 2) | (wt%) (option 1)
1 | 0.977 ± 0.007 | 3 ± 1 | 5 ± 4 | ? | 1.42 ± 0.03
2 | 0.968 ± 0.007 | 5 ± 2 | 7 ± 4 | 2.5 ± 0.5 | 2.84 ± 0.04
3 | 0.926 ± 0.006 | 14 ± 3 | 19 ± 4 | 4.6 ± 0.7 | 5.11 ± 0.07
4 | 0.917 ± 0.006 | 16 ± 4 | 22 ± 4 | 7.9 ± 0.6 | 6.88 ± 0.09
5 | 0.904 ± 0.007 | 21 ± 5 | 25 ± 5 | 9.7 ± 0.5 | 8.7 ± 0.1
No. |  | Molality (mol/kg) |  |  | 
1 | 0.977 ± 0.007 | 3 ± 1 | 5 ± 4 | ? | 1.42 ± 0.03
2 | 0.968 ± 0.007 | 5 ± 2 | 7 ± 4 | 2.5 ± 0.5 | 2.84 ± 0.04
3 | 0.926 ± 0.006 | 14 ± 3 | 19 ± 4 | 4.6 ± 0.7 | 5.11 ± 0.07
4 | 0.917 ± 0.006 | 16 ± 4 | 22 ± 4 | 7.9 ± 0.6 | 6.88 ± 0.09
5 | 0.904 ± 0.007 | 21 ± 5 | 25 ± 5 | 9.7 ± 0.5 | 8.7 ± 0.1
Is m in PTm inputs for SeaFreeze ‘molality’?

## Slide 6
Ice specific heat

## Slide 7
Ice specific heat
Change mass to fit literature ~0.1g (scale mass uncertainty 0.001g)
Propagated uncertainty, literature still falls outside

## Slide 8
Ice specific heat
Effect of using experimental Ice Cp vs literature
~0.2 lower

## Slide 9
Ice specific heat
Is it concerning that Cp is different between our two experiments?
Is it concerning that for 0.1 K/min, there is an increase in datapoint variability?
A quick thought: check Cp from SeaFreeze

## Slide 10
Recalculate Cp using experimental Ice Cp + wt% based on liquidus alignment (option 2)
  For the 1st run, 1.42 wt% was still used because liquidus is above 273.15 K.
Reiterate: in melt phase, the mass fraction of the liquid at the same Temperature is theoretically equal across all samples. Hence we should be expecting same/similar specific heat values.
Recalculations

## Slide 11
Recalculations
Convergence of data (somewhat?)

## Slide 12
Recalculations
Effect of varying wt% (different samples)

## Slide 13
Recalculations
Effect of varying cutoff region

## Slide 14
Tillner-Roth and Friend / SeaFreeze EOS
Comparing with EOS

## Slide 15
Tillner-Roth and Friend / SeaFreeze EOS (Pure Phase)
Comparing with EOS

## Slide 16
At low wt%:
  Greater spread of data
  Greater sensitivity to wt%
Datapoint cut-off determines spline shape
  Spline fit of the whole thermogram is not perfect: sometimes the max/min point is significantly out of line with that of the data
Moving forward, accurately determining wt% is a priority.
Method in choosing cut-off points would also be important.
If we expect the data of different samples to be similar, it can serve as a guiding constraint.
Recalculations (conclusions)

## Slide 17
Moving forward, accurately determining wt% is a priority.
  10 times sample mass measurement: determine errors due to preparation procedure
  Caltech NMR: independent measure of the supply bottle wt%
Recalculations (conclusions)

## Slide 18
77 – 320 K, 0.1 K/min
  20 wt% (done)
  Blank
  15 wt%
Periodicity extraction
Next steps

## Slide 19
Mass of NH4OH (mass balance) = 0.7889 ± 0.001 g
  Some mass fluctuation at 0.0001 scale, thus I manually raised the uncertainty. (It increases density uncertainty by 0.001.
Volume of NH4OH (pipette) = 1.000 ± 0.006 ml
Density of NH4OH = 0.789 ± 0.006 g/ml
Mass Fraction of NH3 = 1.42 – 8.7 ± 0.02 – 0.1 %
Molality of NH3 = 0.85 – 5.6 ± 0.03 – 0.2
Temperature of Sample Error = 0.5 K (Muñoz-Iglesias et. al. 2018)
Error Propagation

## Slide 20
NH3 EOS Errors (specific heat)
I assumed m in PTm inputs in Seafreeze corresponds to molality (mol/kg)
Rotatable .fig files attached in email
Are the following plots what is asked for?

## Slide 21
NH3 EOS Errors (specific heat)
Zoomed in

## Slide 22
NH3 EOS Errors (density)
I assumed m in PTm inputs in Seafreeze corresponds to molality (mol/kg)
Rotatable .fig files attached in email

## Slide 23
Moving averages & splines
Apparent periodicity in 1K bin moving average
Time and T are linearly proportional (interchangeable)

## Slide 24
Moving averages & splines
Periodicity appears in heatflow (both in uncorrected and corrected)
Currently figuring out how to extract frequencies (FFT)

## Slide 25
ADH Double peak
Seem to be an inverse function of amount of NH3 present