## Slide 1
At 0.1K/min, each run takes 3 days.
Conducted blank and 0wt%
8wt% finishes Friday
A pretty short meeting
Summary

## Slide 2
0wt% went smoothly and as expected
Onset T within 1K of expected melting point
0wt% run

## Slide 3
2.5mL vs 5mL
0.1K/min closer to melting point
Will check onset for 0.25K/min on Friday
Visually, it is higher (~275K)
0wt% comparison
Ramp Rate | Onset T | Offset | Peak T | Offset
0.1 | 273.72 | 0.57 | 277.36 | 4.21
0.25 | - | - | 282.20 | 9.05

## Slide 4
Quite a bit lower, somewhat follows Giauque’s trend (lower X)
For Tillner-Roth and Friend’s use case, Zinner provided a better fit for their EoS.
Quick molar specific heat comparison
Tillner-Roth & Friend, 1998

## Slide 5
Total energy absorbed (integral of heat flow) a linear combination of:
Latent heat (ice) * mass of the crystal fraction that have melted in time step
Specific heat of crystal (ice) * mass of crystal fraction present * change in temperature over time step
Specific heat of liquid * mass of liquid fraction present * change in temperature over time step
Specific Heat Retrieval

## Slide 6
Specific Heat Retrieval
(HF1, t1, T1, Fc1,)
(HF2, t2, T2, Fc2)
HF
t
Total Heat:
(HF1+ HF2 )*(t2 - t1) / 2
Melt Ice:
Heat Ice:
based off Feistel & Wagner 2006
Heat Liquid:

## Slide 7
16-19Aug  8 wt%
22-25Aug  1.76 wt%
26-29Aug  3.52 wt%
30-2Sep  6 wt%
5-8Sep  10 wt%
We stick with 1.76 wt% and 3.52 wt% for easy comparison with previous runs.
Experiments to run